## Title
Medical Physics and Clinical Engineering |  |  |  | Addenbrookes Hospital
Clinical Engineering / R&D
Risk File |  |  |  | CE-ER-FRM-22
|  | Product Identification | DELVE Visor
Document Q-Pulse reference
|  | Doc. Name | Risk File - Delve Visor
|  | Mandatory Supporting Documents | Q-Pulse reference
Risk Management Process
Risk Plan
Risk Report
Product description and intended use
Transparent visor covering the face, held in place by an elastic band, and cushioned on the forehead by a foam strip. The visor is considered to be PPE for use by clinical staff in close contact with patients who may have or definitely have COVID-19. The visor: - prevents droplets containing virus being projected onto the skin and mucous membranes of the user's face - reduces probability that direct contact is made with the face by patients with COVID-19 - prevents an asymptomatic, but infected, user projecting particles onto patients. The visor does not prevent small aerosol particles that come around the edges of the visor being entrained in the user's in-breath, and is therefore not considered to provide complete protection and shoudl be usefd tgeter with other appropriate PPE.
Cambridge University Hospitals NHS Foundation Trust |  |  |  |  | Page 1 of 4
This document complies with the Cambridge University Hospitals NHS Foundation Trust service Equality and Diversity statement and must be read in conjunction with Addenbrookes Hospital Policies & Procedures. It is your responsibility to check against the electronic library that this printed out copy is the most recent issue of this document. All changes required should be submitted via Q-Pulse Document module (Change request).
Filename: 008_Risk Evaluation Table
Document No.: CE-ER-FRM-15 |  |  | Version No.: 5 | Active on Q-Pulse & Published

## Revision Table
Document History
Date | Identification | Description | Name
26/03/2020 | 1.0 | Created | AB

## Risk Evaluation Criteria
CUH Risk management strategy and policy 2019-2019
CUH Risk management strategy and policy 2019-2020
Source | CUH Risk management strategy and policy 2019-2020, http://merlin/Lists/DMSRecords/DispRecordTabsDoc.aspx?ID=18427
Adaptations | The risk scoring matrix and grading system has been adapted to enable the CEI team to better track and manage risks that are yet to be determined, or cannot be determined.
Notes | Risk evaluation is for injury or damage to the health of people, or damage to property or the environment, and does include risks surrounding data security. It should NOT consider project or financial impacts. These should instead be considered in the Project Viability Assessment.
Risk Evaluation
| Severity Score |  | Likelihood Score | [6 empty] | Grading Risk
|  |  | 1 | 2 | 3 | 4 | 5 | TBD | CBD | 1-3 | Low risk | No action required
|  |  | Rare | Unlikely | Possible | Likely | Almost Certain | To be determined | Cannot be determined | 4-6 | Moderate risk | Risk reduction if practicable
| 1 | Negligible | 1 | 2 | 3 | 4 | 5 | 25 | 25 | 8-12 | High risk | Risk reduction if practicable, cost benefit analysis required for risks that cannot be practibably lowered
| 2 | Minor | 2 | 4 | 6 | 8 | 10 | 25 | 25 | 15-25 | Significant risk | Unnacceptable risk
| 3 | Moderate | 3 | 6 | 9 | 12 | 15 | 25 | 25 | TBD | To be determined | Activities required to define risk likelihood or severity
| 4 | Major | 4 | 8 | 12 | 16 | 20 | 25 | 25 | CBD | Can't be determined | Treat as unnacceptable and define control measures, or justify lower risk score
| 5 | Catastrophic | 5 | 10 | 15 | 20 | 25 | 25 | 25
| TBD | To be determined | 25 | 25 | 25 | 25 | 25 | 25 | 25
| CBD | Cannot be determined | 25 | 25 | 25 | 25 | 25 | 25 | 25
Likelihood Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Rare | Unlikely | Possible | Likely | Almost Certain
| Frequency | This will probably never happen/recur | Do not expect it to happen/recur but it is possible | Might happen or recur occasionally | Will probably happen/recur but it is not a persisting issue | Will undoubtedly happen/recur,possibly frequently
Severity Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Negligible | Minor | Moderate | Major | Catastrophic
| Impact on the safety of patients, staff or public (physical/psychological harm) | Minimal injury requiring no/minimal intervention or treatment. | Minor injury or illness, requiring minor intervention. Increase in length of hospital stay by 1-3 days Requiring time off work for >3 days | Moderate injury requiring professional intervention. Requiring time off work for 4-14 days Increase in length of hospital stay by 4-15 days RIDDOR/agency reportable incident An event which impacts on a small number of patients. | Major injury leading to long-term incapacity/disability. Increase in length of hospital stay by >15 days. Mismanagement of patient care with longterm effects. Requiring time off work for >14 days | Incident leading to death. Multiple permanent injuries or irreversible health effects. An event which impacts on a large number of patients.
| Quality of treatment | Peripheral element of treatment or service suboptimal Informal complaint/inquiry | Overall treatment or service suboptimal Single failure to meet internal standards | Treatment or service has significantly reduced in effectiveness Repeated failure to meet internal standards | Non-compliance with national standards with significant risk to patients if unresolved Multiple complaints/ independent review Low performance rating Critical report | Totally unacceptable level or quality of treatment/service Gross failure of patient safety if findings not acted on Inquest/ombudsman inquiry Gross failure to meet national standards
| Personal information | No risk to confidential staff or patient data | Occassional/one-off cases of personal data exposure within CUH | Data exposure within CUH | Public exposure of personal data Exposure of personal data to entities who may use that data against the best interest of the person to whom the data belongs | Large scale public exposure of personal data Exposure of personal data to entities who will use that data against the best interest of the person to whom the data belongs
| Envrionmental impact | Minimal or no impact on the environment | Minor impact on environment | Moderate impact on environment | Major impact on environment | Catastrophic impact on environment
| Environmental impact | Minimal or no impact on the environment | Minor impact on environment | Moderate impact on environment | Major impact on environment | Catastrophic impact on environment

## Benefit Evaluation Criteria
008_Risk Evaluation Table
Benefit Evaluation Criteria
Source | Developed by CEI
Adaptations | n/a
Notes | Benefit evaluation is for assessing impact on positive impact on clinical outcome, the patient’s quality of life, outcomes related to diagnosis, positive impact from diagnostic devices on clinical outcomes, or positive impact on public health or patient management. It should NOT consider project or financial impacts. These should instead be considered in the Project Viability Assessment.
Risk Evaluation
| Benefit Size Score |  | Certainty Score |  |  |  |  | Grading Risk
|  |  | 1 | 2 | 3 | 4 | 5 | 15-25 | Large, clear benefit
|  |  | Very unlikely to be realised | Unlikely | Possible | Likely | Almost Certain | 2020-10-12 | Reasonable benefit
| 1 | Negligible | 1 | 2 | 3 | 4 | 5 | 2020-05-09 | Small benefit
| 2 | Minor | 2 | 4 | 6 | 8 | 10 | 2020-01-04 | Very little benefit
| 3 | Moderate | 3 | 6 | 9 | 12 | 15
| 4 | High | 4 | 8 | 12 | 16 | 20
| 5 | Very High | 5 | 10 | 15 | 20 | 25
Likelihood Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Rare | Unlikely | Possible | Likely | Almost Certain
| Frequency | Possible but quite unlikely outcome of the project/extremely uncertain outcome of the project | Possible outcome of the project, but uncertain | Possible outcome of the project | A likely outcome of the project | Will undoubtedly happen as an outcome of the project if completed successfully
Impact Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Negligible | Minor | Moderate | High | Very High
| Medical Benefit | No measurable impact on medical outcomes No measurable impact on patient experience Minor anecdotal benefit | Minor measurable impact on medical outcomes Minor measurable impact on patient experience Moderate anecdotal benefit | Moderate measurable impact on medical outcomes Moderate measurable impact on patient experience Large anecdotal benefit (e.g. patients happy to engage with a home treatment method where they otherwise wouldn't) | Significant measurable impact on key medical outcomes Significant measurable impact on patient experience | Measurable life-saving impact
| Scale of impact | Hospital, short term or one-off Less than 100 patients per year | Hospital Hundreds of patients per year | Hospital or regional Thousands of patients per year | National 10s of thousands of patients per year | International 100s of thousands of patients affected per year
| Clinical benefit | No measurable impact | Minor increase in number of patients treated for a given resource (e.g. reduced hospital stay Minor decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Minor improvement in service quality (e.g. speed of diagnosis, error reduction) | Moderate increase in number of patients treated for a given resource (e.g. reduced hospital stay Moderate decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Moderate improvement in service quality (e.g. speed of diagnosis, error reduction) | Large increase in number of patients treated for a given resource (e.g. reduced hospital stay Large decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Large improvement in service quality (e.g. speed of diagnosis, error reduction) | Transformative increase in number of patients treated for a given resource (e.g. reduced hospital stay Transformative decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Transformative improvement in service quality (e.g. speed of diagnosis, error reduction)

## Overall Benefit Risk
008_Risk Evaluation Table
Benefit Risk Summary
Source | Developed by CEI
| Overall Benefit: | 25
| Residual Risk: | 10
| Summary of benefit: | This device would enable Addenbrookes staff to access face visors when they otherwise would not have access. The project will supply 100,000 disposable face visors, where otherwse the trust would have none. Face visors are a key piece of PPE required to protect staff from liquid droplets transmitting COVID-19 from impacting mucous membranes and skin of the face.
| Summary of risk: | There is a small residual risk that COVID-19 could be transmitted from manufacturer to staff, due to the short timeframe between manufacture and use. COVID-19 can survive on plastic surfaces fro at least 3 days, potentially longer. There is a small risk that someone with a severe latex allergy may be severely harmed through use of this product.
| Actions summary: | If applicable summarise intended actions (for example, to reduce the risk)
Optional: Benefit-risk summary for current state-of-the-art
Source | If the device being assessed is intended to replace/improve upon the current state of the art, it can strengthen the benefit-risk analysis to compare it to the benefit-risk profile of the device representing the current state-of-the-art or best practice. In situations where there is no obvious comparative device, this type of analysis may not be appropriate. “State-of-the-art” embodies what is currently and generally accepted as good practice in technology and medicine. The state of the art does not necessarily imply the most technologically advanced solution.
| Overall Benefit: | 0
| Residual Risk: | 0
| Summary of benefit: | Summarise the overall clinical/medical benefit being delivered
| Summary of risk: | Summarise the overall clinical/medical risk being considered
| Actions summary: | If applicable summarise intended actions (for example, to reduce the risk)
NOT APPLICABLE - product proposal is due to absence of availability of state-of-the-art

## Device Benefit Evaluation
008_Risk Evaluation Table
Benefit Evaluation Table for medical and clinical benefits associated with the medical device
Source | Developed by CEI
| Overall Benefit: | 25 | Evaluate overall residual medical/clinical benefit. This can be estimated as the largest single RPN score.
S = severity, L = likelihood, RPN = Risk Priority Number
Unique ID for hazard plus control measure - useful for referencing hazards in other documents | Brief note explaining benefit | Brief note explaining consequence of benefit | Define any investigatory activities/follow up actions when it is too early to define impact/likelihood | Impact | Likelihood | Score
ID | Benefit | Consequence | Investigatory activities required | S | L | RPN
FV01 | Face visors are available for staff requiring them. Without this project, staff will go without face visors. | Reduction of staff risk in contracting COVID-19 |  | 5 | 5 | 25
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0

## Device Risk Evaluation
Risk Evaluation Table for health hazards associated with the medical device that may impact users or patients
Scope: | injury or damage to the health of people, or damage to property or the environment
Source | Developed by CEI
| Residual Risk: | 10 | Evaluate overall residual medical/clinical risk following the process described in the Risk Management Process CE-ER-PRO-14
| Explanation: | Severity of overall risk is high, as COVID-19 could be transmitted to frontline medical staff via the PPE. However, the risks are being managed through a number of measures to make this unlikely.
S = severity, L=likelihood
Unique ID for hazard plus control measure | Brief note describing hazard | Brief note expalining impact/consequences of hazard | Severity | Likelihood | Score | Define any investigatory activities/follow up actions when it is too early to define a risk control activity | Define activities to control risk (or summarise existing evidence) | Provide a Q-Pulse reference, or location reference to the Technical File | Provide a date of completion for risk control activitie, to confim the activity has been completed | Severity | Likelihood | Score | Detail outcome - is the risk acceptable? If not, is a further risk control measure required (please enter the same hazard with a different ID and repeat the risk evaluation and control prcoess)? Alternatively, is a risk benefit analysis required? | Optional - tick if a risk/benefit is required | Please summarise the outcome of the risk/benefit analysis | If not in this document, please supply a Q-Pulse reference or location reference to the Technicla Filefor location of risk-benefit analysis
ID | Hazard | Consequence | S | L | RPN | Further activities required | Risk Control Measures | Supporting Documentation | Date carried out dd/mm/yyyy | S2 | L2 | RPN2 | Outcome | Risk/Benefit required? | Risk Benefit Outcome | Supporting Documentation2
FV01 | Viral transmission from manufacturer to user via persistence of virus on plastic visor materials | Infection of clinical staff member, who may infect other staff or patients - worst possible consequence is death | 5 | 3 | 15 | Check longevity of viable virus in literature | Process control to ensure risk of transmission from manufacturer to product is minimised | Manufacturing SOP | 29/03/2020 | 5 | 1 | 5 | No further risk reduction practicable for this design. One further risk reduction option is to identify a design that is suitable for cleaning/sterilisation on arrival on site. | Yes | Likelihood of infection through not wearing a visor is probable. The likelihood of infection from a visor subject to these process controls would be extremely rare.
FV02 | Viral transmission due to presence of virus on plastic material for visor before arrival at the manufacturing facility. Transmission to manufacturing staff, or to clinical staff. | Infection of clinical staff member, who may infect other staff or patients - worst possible consequence is death | 5 | 3 | 15 | Risk assess different material supply options | Prevention through material handling controls | Manufacturing SOP | 29/03/2021 | 5 | 1 | 5 | No further risk reduction practicable for this design. One further risk reduction option is to identify a design that is suitable for cleaning/sterilisation on arrival on site. | Yes | Likelihood of infection through not wearing a visor is probable. The likelihood of infection from a visor subject to these process controls would be extremely rare.
FV03 | Viral transmission due to presence of virus on foam material before arrival at the manufacturing facility. Transmission to manufacturing staff, or to clinical staff. | Infection of clinical staff member, who may infect other staff or patients - worst possible consequence is death | 5 | 3 | 15 | Risk assess different material supply options | Prevention through material handling controls | Manufacturing SOP | 29/03/2022 | 5 | 1 | 5 | No further risk reduction practicable for this design. One further risk reduction option is to identify a design that is suitable for cleaning/sterilisation on arrival on site. | Yes | Likelihood of infection through not wearing a visor is probable. The likelihood of infection from a visor subject to these process controls would be extremely rare.
FV04 | Viral transmission due to presence of virus on elastic material before arrival at the manufacturing facility. Transmission to manufacturing staff, or to clinical staff. | Infection of clinical staff member, who may infect other staff or patients - worst possible consequence is death | 5 | 3 | 15 | Risk assess different material supply options | Prevention through material handling controls | Manufacturing SOP | 29/03/2023 | 5 | 1 | 5 | No further risk reduction practicable for this design. One further risk reduction option is to identify a design that is suitable for cleaning/sterilisation on arrival on site. | Yes | Likelihood of infection through not wearing a visor is probable. The likelihood of infection from a visor subject to these process controls would be extremely rare.
FV05 | Viral transmission due to presence of virus on packaging material before arrival at the manufacturing facility. Transmission to manufacturing staff, or to clinical staff. | Infection of clinical staff member, who may infect other staff or patients - worst possible consequence is death | 5 | 3 | 15 | Risk assess different material supply options | Prevention through material handling controls | Manufacturing SOP | 29/03/2024 | 5 | 1 | 5 | No further risk reduction practicable for this design. One further risk reduction option is to identify a design that is suitable for cleaning/sterilisation on arrival on site. | Yes | Likelihood of infection through not wearing a visor is probable. The likelihood of infection from a visor subject to these process controls would be extremely rare.
FV06 | Foam irritates skin | Skin damage | 2 | 3 | 6 |  | Ensure material selected is appropriate for skin contact up to 4 hr | Data sheet | 2020-01-04 | 1 | 1 | 1 | Acceptable | No
FV07 | Foam irritates skin | clinical staff stop wearing mask | 5 | 2 | 10 |  | Ensure material selected is appropriate for skin contact up to 4 hrs | Data sheet | 2020-01-04 | 5 | 1 | 5 | No further risk reduction practicable | Yes | Likelihood of infection through not wearing a visor is highly likely. The likelihood of infection from a visor that is removed due to discomfort is also highly likely. Having the mask does not worsen the risk profile compared to the baseline situation.
FV08 | Droplets impact facial skin despite wearing visor | Infection, death | 5 | 2 | 10 |  | Dimensions of mask match or exceed predicate device - height above face, below face, width, extent of curvature | Design verification & validation | 2020-01-04 | 5 | 1 | 5 | No further risk reduction practicable | Yes | This design of mask provides the same or better protection than CE marked designs.
FV09 | Mask fogs up during use | clinical staff stop wearing mask | 5 | 2 | 10 | Evaluate need for anti-fog | Testing for 4hrs whilst wearing FFP3 mask | Design verification & validation | 2020-01-04 | 5 | 1 | 5 | No further risk reduction practicable | Yes | Likelihood of infection through not wearing a visor is highly likely. The likelihood of infection from a visor that is removed due to discomfort is also highly likely. Having the mask does not worsen the risk profile compared to the baseline situation.
FV10 | Mask does not fit face properly | clinical staff stop wearing mask | 5 | 2 | 10 | Initial evaluation by range of staff required - could we put a QR code on visor to a GoogleForm? | Testing with a range of different head size users | Design verification & validation | 2020-01-04 | 5 | 1 | 5 | No further risk reduction practicable | Yes | Likelihood of infection through not wearing a visor is highly likely. The likelihood of infection from a visor that is removed due to discomfort is also highly likely. Having the mask does not worsen the risk profile compared to the baseline situation.
FV12 | Mask is uncomfortable by design | clinical staff stop wearing mask | 5 | 2 | 10 | Initial evaluation by range of staff required - could we put a QR code on visor to a GoogleForm? | Testing for 4hrs whilst wearing FFP3 mask | Design verification & validation | 2020-01-04 | 5 | 1 | 5 | No further risk reduction practicable | Yes | Likelihood of infection through not wearing a visor is highly likely. The likelihood of infection from a visor that is removed due to discomfort is also highly likely. Having the mask does not worsen the risk profile compared to the baseline situation.
FV13 | Mask is uncomfortable due to sharp edges or staples | clinical staff stop wearing mask | 5 | 2 | 10 |  | Quality control during manufacturing process | Manufacturing SOP | 29/03/2024 | 5 | 1 | 5 | No further risk reduction practicable | Yes | Likelihood of infection through not wearing a visor is highly likely. The likelihood of infection from a visor that is removed due to discomfort is also highly likely. Having the mask does not worsen the risk profile compared to the baseline situation.
FV14 | Non-viral infection of clinical staff by face sheild | Staff sickness - possibility of passing on to patient | 4 | 2 | 8 | Risk assess different material supply options | Prevention through material selction/storage controls | Manufacturing SOP | 29/03/2024 | 4 | 1 | 4 | Acceptable | No
FV15 | Elastic contains latex | Skin damage | 2 | 2 | 4 |  | First, try to source latex free elastic. If not possible, mitigate through selecting elastic where the latex material is enclosed thereby unlikely to come into contact with skin, and clearly label the product stating that the elastic contains latex. | Data sheet Product Labels | 2020-01-04 | 2 | 1 | 2 | Acceptable | No
FV16 | Elastic contains latex | irritates skin, clinical staff stop wearing mask | 5 | 1 | 5 |  | First, try to source latex free elastic. If not possible, mitigate through selecting elastic where the latex material is enclosed thereby unlikely to come into contact with skin, and clearly label the product stating that the elastic contains latex. | Data sheet Product Labels | 2020-01-04 | 5 | 1 | 5 | No further risk reduction practicable | Yes | Likelihood of infection through not wearing a visor is highly likely. The likelihood of infection from a visor that is removed due to discomfort is also highly likely. Having the mask does not worsen the risk profile compared to the baseline situation.
FV17 | Elastic contains latex | Severe allergic reaction | 5 | 1 | 5 |  | First, try to source latex free elastic. If not possible, mitigate through selecting elastic where the latex material is enclosed thereby unlikely to come into contact with skin, and clearly label the product stating that the elastic contains latex. | Data sheet Product Labels | 2020-01-05 | 5 | 1 | 5 | No further risk reduction practicable | Yes | Severe allergic reactions to latex are extremely rare. The likelihood of infection of any member of staff through not wearing a visor is very high.
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0

## PHA
| Risk Questions | N/A | Detail
1.2 | Is the PPE intended to contact the patient or other persons? | Yes | Contact for up to 4 hours with hair and skin around forehead
1.3.1 | Are all materials used in the PPE biocompatible? | TBD | Materials have not been selected for biocompatibility, but are considerd to be low risk
1.3.2 | Will any harmful materials or substances be in contact with the patient? | No
1.4 | Is energy delivered to and/or extracted from the patient? | No
1.5 | Are substances delivered to and/ or extracted from the patient? | No
1.6 | Are biological materials processed by the PPE for subsequent re-use? | No
1.7.1 | Is the PPE supplied sterile or intended to be sterilised by the user? | No | Meaures are taken to minimise risk of virus being present on the visor, but sterility cannot be guaranteed
1.7.2 | Are other microbiological controls applicable?
1.8 | Is the PPE intended to be routinely cleaned and disinfected by the user? | TBD | Re-cleaning after use is under investigation. In first instance the PPE will be assumed to be single use.
1.9 | Is the PPE intended to modify the patient environment? | No
1.1 | Are measurements taken by the device? | N/A
1.11 | Is the PPE interpretative? | N/A
1.12.1 | Is the PPE intended for use in conjunction with medicines? | N/A
1.12.2 | Is the PPE intended for use in conjunction with other medical technologies? | N/A
1.13 | Is there any unwanted output of energy or substances? | No
1.14 | Is the PPE susceptible to environmental influences? | No
1.15 | Does the PPE influence the environment? | No
1.16 | Are there essential consumables or accessories associated with the PPE? | No
1.17.1 | Is maintenance necessary? | No | To be single use or disposed of in event of failure to perform adequately.
1.17.2 | Is calibration necessary? | No
1.18 | Does the PPE contain software? | No
1.19 | Does the PPE have a restricted shelf-life? | TBD
1.2 | Are there any delayed and/or long-term use effects? | N/A
1.21 | Will the PPE be subject to mechanical forces? | No | No forces beyond routine operation - putting on and taking off.
1.22.1 | Does the PPE have a battery that will affect the lifetime of the device? | No
1.22.2 | Will ageing determine the lifetime of the PPE? | No
1.23 | Is the PPE intended for single use? | Yes | Yes, pending definition of an acceptable cleaning regime
1.24 | Is safe decommissioning or disposal of the PPE necessary? | Yes | Clinical waste after use.
1.25.1 | Does installation of the PPE require special training? | N/A
1.25.2 | Does use of the PPE require special training? | No | Intuitive to use.
1.26 | Will new manufacturing processes need to be established or introduced? | Yes | See SOP created for manufacturing
1.27 | Is successful application of the PPE critically dependant on human factors? | No
1.27.1 | Does the PPE have connecting parts or accessories? | No
1.27.2 | Does the PPE have a control interface? | No
1.27.3 | Does the PPE display information? | No
1.27.4 | Is the PPE controlled by a menu? | No
1.28 | Is the PPE intended to be mobile or portable? | Yes | By design

## Classification
The Face Visor is classified as PPE under Regulation (EU) 2016/425 on personal protective equipment and must meet the requirements of Annex II of the Directive.
The Face Visor is not a medical device.
https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX:32016R0425
The Face Visor is Category 3 according to the PPE Reg since its failure to perform as required may expose the user inadvertently to harmful biological agents, i.e. COVID-19 virus
RISK CATEGORIES OF PPE
Category I
Category I includes exclusively the following minimal risks:
(a) | superficial mechanical injury;
(b) | contact with cleaning materials of weak action or prolonged contact with water;
(c) | contact with hot surfaces not exceeding 50 °C;
(d) | damage to the eyes due to exposure to sunlight (other than during observation of the sun);
(e) | atmospheric conditions that are not of an extreme nature.
Category II
Category II includes risks other than those listed in Categories I and III;
Category III
Category III includes exclusively the risks that may cause very serious consequences such as death or irreversible damage to health relating to the following:
(a) | substances and mixtures which are hazardous to health;
(b) | atmospheres with oxygen deficiency;
(c) | harmful biological agents;
(d) | ionising radiation;
(e) | high-temperature environments the effects of which are comparable to those of an air temperature of at least 100 °C;
(f) | low-temperature environments the effects of which are comparable to those of an air temperature of – 50 °C or less;
(g) | falling from a height;
(h) | electric shock and live working;
(i) | drowning;
(j) | cuts by hand-held chainsaws;
(k) | high-pressure jets;
(l) | bullet wounds or knife stabs;
(m) | harmful noise.